ANNEX ΙΙ: TERMS OF REFERENCE – TECHNICAL SPECIFICATIONS 
Table of Contents

1.	CONTRACT SCOPE	2

1.1	Requirements analysis – Specifications	2

1.2	Reporting requirements	3

1.3	Description of Additional Requirements/ Services	3

2.	FACILITIES TO BE PROVIDED BY THE CONTRACTING AUTHORITY	4

3.	DELIVERY – ACCEPTANCE OF CONTRACT SCOPE	4

3.1	Place of delivery	4

3.2	Contract Execution schedule	4

3.3	Checks and acceptance of deliverables	5

# CONTRACT SCOPE

## Requirements analysis – Specifications

The present Contract concerns the supply, installation and training for one (1) Inductively Coupled Plasma-Mass Spectrometry (ICP-MS) system, intended for trace and ultra-trace multi-elemental analysis in environmental and related matrices.

The proposed ICP-MS system shall fully comply with the technical specifications listed in Form 9 (Tenderer’s Technical Response and Compliance Table), which forms an integral part of this Contract.

The specifications cover all major sub-systems of the ICP-MS, including: 1) Sample introduction system, which comprises of a peristaltic pump, a Peltier-cooled double-pass Scott‑type quartz spray chamber, a concentric nebulizer and an auto-sampler with a corrosion-resistant enclosure, with spill tray and rinse port and capacity ≥ 240 for 17 mm OD tubes. 2) RF generator, delivering 0.5-1.6 kW power range at 27 MHz frequency or better. 3) Plasma torch and interface, with automatic XYZ alignment, oxide ratio ≤ 0.5% (CeO⁺/Ce) under high‑matrix mode and suitable Ni/Pt cones allowing maintenance/cleaning/replacement without breaking the vacuum. 4) Vacuum and ion-focusing systems, fully software-regulated and easy to maintain, ensuring automatic recovery after power loss and low background. 5) Collision/reaction cell with multipole design, ≥ 2 gas lines, switching time ≤ 5 s, operated entirely through system software. 6) Quadrupole mass analyzer and detector, offering mass range 2-260 u, scan speed> 3000 amu/s, dynamic range ≥ 11 orders and background <1 cps. 7) Data system and software, providing full control of all subsystems, automated optimization, quantitative/semi-quantitative analysis, method-development wizards, nanoparticle/single-cell module and comprehensive QC monitoring. 8) A PC workstation, including specific peripherals. The supply shall include all parts, accessories, consumables, certified reference materials and others required for system installation, start-up and verification of analytical performance.

The supplier shall ensure full CE-marking, RoHS and EMC compliance and provide corresponding certificates and EU Declarations of Conformity for all sub-systems (including the auto-sampler and peripherals).

Preventive maintenance and technical support during the warranty period shall be performed by factory-trained personnel, ensuring reliable performance and minimal downtime. The warranty shall cover parts, labour and travel expenses and its cost is considered included in the total offer price.

The proposed system shall be fully compatible with Liquid Chromatography (LC) and Gas Chromatography (GC) speciation accessories, providing expandability.

All Quantities, Mandatory Requirements and Performance Specifications are detailed in Form 9, to which this Annex explicitly refers.

## Reporting requirements

The Contractor should submit to the Contracting Authority the following Reports on the implementation of Contract Scope activities and the results achieved:

- Inception Report – within 2 weeks from Contract signature.

- Completion Report – at the end of the Contract implementation period.

The Inception Report should contain the following:

- The key issues identified (if any).

- The updated schedule for deliveries, installation activities, trainingand any other activities required under the Contract Scope, indicating the critical points.

The Completion Report should contain the following:

- A list of all implemented deliveries, installation activities and any other activities required, and of any deliverables developed.

- An assessment of the results achieved.

- The Contractor shall submit Ad-hoc Reports whenever it deems that it should inform the Contracting Authority of significant issues.

The Reports should be submitted electronically to the Contracting Authority, for the attention of the Project Manager. The Project Manager is responsible for the submission of the reports to the competent Acceptance Committee.

## Description of Additional Requirements/ Services

- On-site Installation: Installation shall be performed on site by factory-trained personnel certified by the manufacturer, within 1 month from the delivery.

- On-site Users training: A comprehensive training programme (minimum 4-5 trainees) shall be delivered by factory-authorized personnel covering operation and maintenance, method development, tuning and data evaluation, QA/QC and troubleshooting. All manuals (Operation, Maintenance,Installation) shall be provided in English and/or Greek, in electronic and/or printed format. Training is expected to take place following the installation and within maximum 2 months from installation.

# FACILITIES TO BE PROVIDED BY THE CONTRACTING AUTHORITY

• During the on-site visit for the installation of the supplies, the Contractor will be given access to the installation site (The Cyprus Institute, Environmental Chemistry Laboratory, NTL Building, 2nd Floor, Aglantzia).
• In coordination with the Contractor, all necessary on-site support to install the supply will be provided (i.e. IT support, Electrical provisions, Health and Safety aspects).

# DELIVERY – ACCEPTANCE OF CONTRACT SCOPE

## Place of delivery

DELIVERY AND INSTALLATION SITE:

Environmental Chemistry Laboratory (ECL), CARE-C, The Cyprus Institute
POSTAL ADDRESS: 20, Konstantinou Kavafi Str. 2121, Aglantzia
CONTACT PERSON: 00 357 22 208 722 (Minas Iakovides)

## Contract Execution schedule

Delivery must take place no later that 5 months from contract signing.
Installation should take place no later than 1 month from the delivery.
Training should take place no later than 2 months from the delivery following installation.

| Product Delivery | 5 months from the date of Contract signing |
| --- | --- |
| Additional Services:(*) | 3 months from the date of product delivery |
| • Installation/ Setup | 1 month from the date of product delivery |
| • User Training | Maximum 2 months from the date of product installation. |
| Maximum Time for Final Acceptance of the Contract Scope | 8 months from the date of Contract signing |
| Warranty Period | 12 months |
| Maintenance Period | 12 months from the date of final acceptance of the Contract Scope |
| Contract Duration (excluding extension rights period) | 20 months from the date of Contract signing |
| TOTAL CONTRACT DURATION (maximum) | 20 months from the date of Contract signing |

## Checks and acceptance of deliverables

The acceptance of the requested ICP-MS system and all associated components will be carried out by the Acceptance Committee of the Contracting Authority following delivery, installation,  QA/QC and completion of the on-site training and shall take place no later than thirty (30) days after delivery, installation, QA/QC and on-site training. The acceptance procedure will include among others macroscopic inspection, QA/QC performance verification (execution of operational tests confirming proper functioning of the system and all sub-systems). Additionally, analytical performance will be verified via measurements of certified reference materials or equivalent solutions to confirm compliance with the contractual performance criteria.