## FORM 9

# TENDERER’S TECHNICAL RESPONSE AND COMPLIANCE TABLE

### 1. Requested product: One (1) Inductively Coupled Plasma–Mass Spectrometry (ICP-MS) system

| No. | Features – Specifications | Requirement | Response – Offer | Reference |
| --- | --- | --- | --- | --- |
|  | GENERAL FEATURES |  |  |  |
|  | Manufacturer. | YES |  |  |
|  | Model. | YES |  |  |
|  | SPECIFIC FEATURES |  |  |  |
|  | Bench top instrument. | YES |  |  |
|  | Instrument footprint (including the PC, the monitor, keyboard/mouse and the auto-sampler): ≤ 1500 mm (length) × 750 mm (width) × 900 mm (height). | YES |  |  |
|  | CE‑marking and Full RoHS & EMC compliance (Certificates Required). | YES |  |  |
|  | Compatibility with LC and GC speciation kits (hardware & software). | YES |  |  |
|  | The system should feature: - Liquid sample introduction, including the auto-sampler. - Solid-state RF generator and plasma torch. - Plasma/MS interface. - Vacuum pump system. - Ion focusing system. - Collision/reaction cell. - Quadrupole mass analyzer. - Detector. - Appropriate software and hardware for system control and data analysis. | YES |  |  |
|  | Sample‑introduction pathway, enabling routine analysis of aqueous matrices containing ≥ 25% w/w total dissolved solids (TDS) without off‑line dilution and/or torch replacement. | YES |  |  |
|  | SAMPLE INTRODUCTION SYSTEM |  |  |  |
|  | Peristaltic Pump: ≥ 3 channels, low‑pulsation and high precision, adjustable speed (10‑roller). | YES |  |  |
|  | Concentric quartz Nebulizer, nominal uptake ca. 0.2 mL/min. | YES |  |  |
|  | Double‑pass Scott‑type Quartz spray chamber, Peltier‑cooled (-5 °C to +20 °C or better). | YES |  |  |
|  | The sample introduction system must be fully regulated, monitored and safeguarded by the system software. | YES |  |  |
|  | The sample introduction system must be easily accessible to the user, ensuring straightforward repairability-maintenance. | YES |  |  |
|  | AUTOSAMPLER |  |  |  |
|  | Fully user-programmable in x, z and θ dimensions & monitored by the system software. | YES |  |  |
|  | CE marking, EU Declaration of Conformity & Certificates to signify compliance with EMC directives of the European Union. | YES |  |  |
|  | Self diagnostics for operational errors or instrument status | YES |  |  |
|  | Corrosion‑resistant autosampler with spill tray and rinse port, with random access to ≥ 240 for 17 mm OD tubes. | YES |  |  |
|  | Includes an enclosure to protect the instrument and the surrounding environment from corrosion, with extraction air duct. | YES |  |  |
|  | RF GENERATOR |  |  |  |
|  | Solid-state digital drive radio-frequency (RF) generator, with appropriate cooling system. | YES |  |  |
|  | RF frequency 27 MHz or better (if better, details must be given for evaluation). | YES |  |  |
|  | RF Power range from 0.5 to 1.6 KW. | YES |  |  |
|  | RF Generator System must be regulated/monitored by the system software. | YES |  |  |
|  | The RF generator must be maintenance-free. | YES |  |  |
|  | PLASMA |  |  |  |
|  | Automatic X‑Y‑Z torch alignment, step size ≤ 0.1 mm. | YES |  |  |
|  | One-piece Quartz Plasma Torch. | YES |  |  |
|  | The torch fitting should be suitable for fast assembly/disassembly, without excessive alignment. | YES |  |  |
|  | Plasma ignition and optimization must be fully automatic through the start up procedure. | YES |  |  |
|  | Torch & Interface design delivering CeO /Ce oxide ratio ≤ 0.5% under high‑matrix mode. | YES |  |  |
|  | PLASMA / MS INTERFACE |  |  |  |
|  | Plasma/MS interface equipped with suitable Ni and Pt-tipped sampling and skimmer cones. | YES |  |  |
|  | The cones must be easily removed for cleaning, maintenance and/or replacement, without breaking the vacuum of the system. | YES |  |  |
|  | VACUUM SYSTEM |  |  |  |
|  | Vacuum system fully regulated by the system software. | YES |  |  |
|  | Appropriate filtering for the external rotary vacuum pump to prevent oil fumes & odors, easily replaceable for the user. | YES |  |  |
|  | A suitable and distinguishable vacuum system, with split-flow turbo molecular pump and single external rotary pump. | YES |  |  |
|  | Auto vacuum recovery after power loss. | YES |  |  |
|  | ION FOCUSING SYSTEM |  |  |  |
|  | Ion Focusing System easy to handle and maintain by the user. | YES |  |  |
|  | The System must effectively prevent photons and neutral particles from the plasma reaching the detector, giving low random background. | YES |  |  |
|  | COLLISION / DYNAMIC REACTION CELL |  |  |  |
|  | Collision / dynamic reaction cell that minimizes polyatomic mass interferences, designed to operate under various gases. All cell functions must be regulated / monitored by the user, via the system software. | YES |  |  |
|  | Multipole collision / reaction cell, with ≥ 1 dedicated collision gas line and ≥ 1 spare reaction‑gas line, and able to accept a third cell gas line. | YES |  |  |
|  | Cell able to be operated under: i) standard mode: without gas; ii) collision mode: with a non-reactive gas; and upgraded with iii) reaction mode: with a reactive gas. | YES |  |  |
|  | Switching between operation modes must be done through the system software, without interventions from the user in the system hardware. | YES |  |  |
|  | Gas‑line switching time ≤ 5 s. | YES |  |  |
|  | MASS ANALYZER & DETECTOR |  |  |  |
|  | Quadrupole mass filter operating at ≥ 3 MHz, with hyperbolic rods. | YES |  |  |
|  | Mass Resolution from 0.3 u to 1.0 u. | YES |  |  |
|  | Mass Range: 2 u to 260 u, or higher. | YES |  |  |
|  | Mass Stability: ≤ 0.05 u / day; ≤ 0.1 u / six months. | YES |  |  |
|  | Mass Scan Speed >3000 amu / s. | YES |  |  |
|  | Detector assembly providing continuous linear dynamic range ≥ 11 orders (0.1 cps - 10 Gcps), without manual range switching. | YES |  |  |
|  | Minimum dwell time ≤ 0.1 ms and data acquisition speed ≥ 10000 reads / s. | YES |  |  |
|  | Background: < 1 cps. | YES |  |  |
|  | Sensitivity (Mcps/ppm): 7Li: 140, 59Co: 400, 89Y: 600, 115In: 700, 205T: 520, 238U: 720, or better. | YES |  |  |
|  | Abudance sensitivity in low and high mass sides (Cs): ≤ 5E-7 and 1E-7, respectively. | YES |  |  |
|  | PERFORMANCE |  |  |  |
|  | Instrument detection limits (ppt): 9Be: 0.05, 115In: 0.02, 209Bi : 0.02, or better | YES |  |  |
|  | Oxide ratio CeO+/Ce ≤ 1.5% in standard (no gas) mode. | YES |  |  |
|  | Doubly‑charged ions Ce2+/Ce ≤ 2.5% in standard (no-gas) mode. | YES |  |  |
|  | Long-term stability (2 h): ≤ 1.5% RSD. Short-term stability (20 min): ≤ 1% RSD. | YES |  |  |
|  | DATA PROCESSING (SOFTWARE) & DATA QUALITY |  |  |  |
|  | Software provides full instrument control (initiation, configuration, optimization), diagnostics information, maintenance feedback and performance logging of operational events. | YES |  |  |
|  | Software supports full chiller control (fluid level, temperature), error notifications and unattended shut-down. | YES |  |  |
|  | Software provides data exporting to compatible spreadsheet software programs (i.e. Microsoft Excel). | YES |  |  |
|  | Integrated nanoparticle and single‑cell analysis module. | YES |  |  |
|  | Software enables fully automated and unattended system stand-by / shut-down after the termination of a sequence. | YES |  |  |
|  | Software able to perform Quantitative and Semi-Quantitative Analysis. | YES |  |  |
|  | Software able to guide the user through hardware configuration, instrument optimization, sample acquisition, data analysis and report. | YES |  |  |
|  | Software able to provide real-time results, interactive calibration plots, on-screen flags for outliers, internal standard recovery, QC stability plots, mass spectrum, counts, concentration values, RSD of concentration values and data storage. | YES |  |  |
|  | Method development automation, with setup wizards for building fully functional methods for different applications. | YES |  |  |
|  | Software capable of monitoring instrument status and notifying the user upon maintenance based on several metrics (e.g. operation time, number of samples measured, etc). | YES |  |  |
|  | PERSONAL COMPUTER (PC) |  |  |  |
|  | Computer running latest version of Windows. | YES |  |  |
|  | Solid state Disk (SSD) ≥ 1.0 TB. | YES |  |  |
|  | Memory (RAM) ≥ 16 GB. | YES |  |  |
|  | Multi-core CPU performance ≥ 4.00 GHz, compatible with workstation PCs. | YES |  |  |
|  | Monitor ≥ 24 in. | YES |  |  |
|  | USB ports ≥ 6. | YES |  |  |
|  | Keyboard and optical mouse included. | YES |  |  |
|  | INSTALLATION & TRAINING |  |  |  |
|  | System should be supplied complete, including all necessary accessories and consumables needed for installation, system checkout and immediate initiation. | YES |  |  |
|  | Operator, Installation and Maintenance Manuals (in English or Greek) of the entire instrument (including the chiller & auto-sampler) must be supplied upon installation in electronic and/or hard copy form. | YES |  |  |
|  | MAINTENANCE - TECHNICAL SUPPORT |  |  |  |
|  | System maintenance and technical support must be conducted by factory-trained personnel (a corresponding updated certificate is required). | YES |  |  |
|  | Warranty period (including instrument maintenance, labour and maintenance cost, parts cost) at least 12 months. | YES |  |  |

NOTES:

  - Responses MUST be provided for all items in the Table. If no response is provided for a particular item, this shall be considered as a negative response.

  - The Column “FEATURES – SPECIFICATIONS” or “REQUIREMENTS – SPECIFICATIONS” is used to describe in detail the corresponding technical terms, obligations or clarifications to which responses must be provided.

  - If the Column “REQUIREMENT” contains the word “YES” (meaning that the corresponding specification is mandatory for the Tenderer) or a specific number (meaning that the specification has a mandatory numerical value which must be complied with), these shall be considered as unchangeable terms and the Technical Offers that do not meet such requirements shall be rejected as inadmissible.

  - The Tenderer’s response is entered in the Column “RESPONSE – OFFER”. This response is either the specific product that the Tenderer offers or is of the form “YES/NO” (depending on whether the Technical Offer does or does not meet the respective specification) or a specific numerical value denoting the quantity of the respective feature in the Technical Offer.

  - The Column “REFERENCE” is used to enter an unambiguous reference to an Annex of the Technical Offer containing numbered Technical Brochures of manufacturers, or detailed technical descriptions of the products offered, which in the opinion of the Tenderer document the information stated in the Table. References should be specific (e.g. Technical Brochure 3, p. 4, par. 4), while in the technical brochure or report in question the respective point shall be underlined and a note shall be made of the paragraph in the Table where the required feature is stated.